## Slide 1: Polarization Calibration of the Daniel K Inouye Solar Telescope (DKIST) formerly Advanced Technology Solar Telescope
David Elmore
Instrument Scientist
National Solar Observatory1
Boulder, CO USA
1 Association of Universities for Research in Astronomy funded by the National Science Foundation.

## Slide 2
4-m off axis Gregorian
Altitude-Azimuth mount
At least 10 mirrors between the sun and an instrument

## Slide 3
Calibration
Polarization state generation optics at Gregorian focus to calibrate the relay optics and polarimeters.
Polarimeters
Polarimeters at each of four instruments in the Coudé Lab
Use observations to calibrate mirrors preceding calibration optics.

## Slide 4: Coudé Layout
ViSP
Cryo-NIRSP
VTF
DL-NIRSP
VBI
WFC

## Slide 5
Cryogenic Near Infrared Spectro-polarimeter (Cryo-NIRSP): Individually selectable lines between 1000nm and 5000nm.
Diffraction Limited Near Infrared Spectro-polarimeter (DL-NIRSP): Three selectable lines simultaneously between 500nm and 2500nm.
Visible Spectro-polarimeter (ViSP): Any three lines simultaneously between 380nm and 1100nm
ViSP
DL NIRSP
Cryo NIRSP
VTF
Visible Tunable Filter: Individually selectable lines between 520nm and 860nm.

## Slide 6
Create a model, a string of Mueller matrices, that describes polarization of the telescope end-to-end.
Measure intensities for numerous calibration optics generator states and telescope orientations to infer parameters of the model from Gregorian focus to the instruments
Measure polarization of targets of known polarization signatures to infer properties of optics preceding the Gregorian focus
Polarization Calibration Process

## Slide 7
Polarization model predicts:
Intensity for each calibration optical configuration (j), elevation, coude-azimuth angle, and modulator rotation angle (φ) at a range of wavelengths.
At least a starting point for higher order effects, field of view changes, temporal degradation, …
Calibration
Polarimeters

## Slide 8
Group mirrors into fewer individual matrices
  M12
  M34
  M56
Use a simplified matrix description for each of the groups
Group final mirrors, instrument optics, and modulators into a modulation matrix
In order to infer the telescope model from measurements, simplify the model to reduce the number of free parameters

## Slide 9: Mirror Groups
A string of mirrors with axes crossed or aligned has the matrix form of a single mirror and is described by retardation, δ, and the ratio of reflectivities in and perpendicular to the plane of incidence, rs/rp.
M56
M34
M12

## Slide 10
Solid = Azimuth Dotted = Elevation Winter, Spring, Summer

## Slide 11: Modulation Matrix
Modulation matrix includes the modulator itself as well as all polarizing optics from the last telescope rotation through the instrument.
-- José Carlos del Toro Iniesta, Manolo Collados.
Oinst

## Slide 12
Calibration polarization generator configurations.
-- Andrew Skumanich, Bruce Lites, Valentin Martinez Pillet
Clear
Retarder: 0°, 45°, 90°, 135°
Linear polarizer: 0°, 45°, 90°, 135°
Polarizer + retarder: Polarizer 0°, 45°, 90°, 135°,
Retarder = Polarizer ± 45°, 0°
Also determine:
Input Stokes vector
Transmission of retarder.
Transmission of polarizer.
Mount error of retarder relative to polarizer.
Retardation of retarder!

## Slide 13
Truth: solid
Simple: diamonds

## Slide 14: M1 and M2
From observations, one must determine two parameters, δ, and the ratio of reflectivities in and perpendicular to the plane of incidence, rs/rp.

## Slide 15: M1 and M2

## Slide 16
Polarimeters
Figure . Synthetic profiles before and after correcting for Doppler shift. Polarization parameters are magnified by a factor of five. Black traces are original profiles. Colored curves are Doppler corrected I(magenta), Q (red), U(green), and V(blue).
Correlation Method: Zeeman effect in magnetically sensitive absorption lines

## Slide 17
Polarimeters
Circular to linear crosstalk using the product of a circular polarization kernel times Q and times U plotted vs. the kernel times V.
Correlation Method

## Slide 18
Rayleigh scattering: Harrington, Kuhn & Hall 2011

## Slide 19
rs/rp is determined from the measured offset seen 0° (Q)
Rayleigh scattering
δ is determined by the amount of measured circular polarization.
X
X
Measure sky polarization including locations with the scattering angles at ±45° (±U) to the azimuth of the telescope and therefore plane of incidence on M1M2.
X
X

## Slide 20
rs/rp is determined from orientation of linearly polarized light, should be tangent to the limb
Coronal Polarization:
δ is determined by the amount of measured circular polarization.
Rayleigh scattering goes as λ-4 At 4μ the K-corona is brighter than the sky!

## Slide 21
Create an end-to-end polarization model of the telescope and instruments.
Measure intensities for numerous calibration optics generator states and telescope orientations to infer parameters of the model from Gregorian focus through each polarimeter
Perform polarization measurements of targets of known polarization signatures to infer properties of optics preceding the Gregorian focus
DKIST Polarization Calibration Process

## Slide 22
Division of time: High modulation efficiency simultaneously over a 5:1 wavelength range
4k x 4k field of view: Division of wave front a problem.
Diffraction limited 4-m telescope: Division of aperture a problem.
Full Stokes polarimetry at a continuous range of wavelengths in high spectral resolution: Spectral modulation a problem.

## Slide 23
Rotating Poly-Chromatic Modulator designed by Roberto Casini consisting of three quartz compound zero order wave plates at unique angles producing an elliptical retarder.

## Slide 24
Using efficiency methodology published by Jose Carlos del Toro Iniesta and Manolo Collados, sampled over 20 states per rotation (in this example) efficiency vs. wavelength is:
100%
90%
80%
100%
90%
80%
100%
90%
80%