## Slide 1
Mueller polarimetry through an optical fiber
1J. Vizet, 1S. Manhas, 2S. Deby, 2J.C. Vanel, 2A. De Martino, 1D. Pagnoux
1 Institut de recherche XLIM, UMR CNRS 7252,
Université de Limoges, Faculté des Sciences et Techniques, Limoges, France
2 LPICM, UMR CNRS 7647,
Ecole Polytechnique, Palaiseau, France

## Slide 2
Why polarimetric imaging of biological tissues ?
Mueller polarimetry
H
A
Colon cancer [1]
Depolarization
Cervix cancer [2]
Retardance
H : Healthy , A : Abnormal
A
H
H
M : 4x4 Mueller matrix
Depolarization
Diattenuation
(linear and circular)
Retardance
(linear and circular)
A
[1] : “Ex-vivo characterization of human colon cancer by Mueller polarimetric imaging”. Angelo Pierangelo et al., Optics express 1593, Vol. 2, No. 9 (2011)
[2] : “Imagerie polarimétrique pour le diagnostic du cancer du col utérin” A. Pierangelo et al., Journées d’imagerie non-conventionnelle (2013)
Applied to biological tissues imaging …

## Slide 3
PSG
PSA
Biological
Sample
Imaging
reconstruction
system
Polarimetric imaging of biological tissues
PSG : Polarization states generator
PSA : Polarization states analyzer
Probing
Polarization
State
Backscattered
Polarization state
Light source
Detector
DIRECTLY ANALYSED
WELL KNOWN
Drawbacks :
Need of biopsies
Time consuming
Problematic of the use of Mueller imaging systems
Polarimetric image
of biological sample

## Slide 4
Why polarimetric imaging of biological tissues through an endoscopic fiber ?
Light source
Image reconstruction system
Optical fiber (endoscope)
Polarization analysis system
Detector
Trachea
Bronchi
Abnormal tissue
Advantages :
In vivo in situ observations
Possibility of early detection of diseases
Less biopsies
PSG
PSA

## Slide 5
Why polarimetric imaging of biological tissues through an endoscopic fiber ?
Problem :
Optical fiber modifies polarization states in a uncontrolled manner
Probing AND backscattered states are UNKNOWN
Optical fiber (endoscope)
Image reconstruction system
Trachea
Bronchi
Abnormal tissue
Light source
Polarization analysis system
Detector
PSG
PSA

## Slide 6
Existing techniques for polarimetric endoscopic characterizations
[3] : “Fiber-optic device for endoscopic polarization imaging”. J. Desroches et al, Opt. Lett. 34, 3409-3411 (2009)
[4] : “Depolarization Remote Sensing by Orthogonality Breaking” . J. Fade and M. Alouini, PRL 109, 043901 (2012)
Advantages :
Measurement insensitive to propagation in fiber
No specific component is needed near sample
Orthogonality breaking of two waves coming from a single laser [4] :
Drawback :
Both depolarization and diattenuation
cause orthogonality breaking
Compensation of fiber birefringence by the use of a Faraday rotator [3] :
Advantages :
Linear retardance measurement of samples
Drawback :
Measurement can be slow

## Slide 7
[5] : “Narrow band 3x3 Mueller polarimetric endoscopy”. Ji Qo et al, Opt. Express, 14, 2433-2449 (2013)
Polarimetric analysis through a rigid laparoscope [5] :
Rat abdomen
Raw image
Depolarization
image
Advantages :
Large field of view (5,5 x 5,5cm)
Avoid complicated miniaturizations
Drawbacks :
PSG states generated by rotation of the laparoscope
Spatial stability problems
3x3 Mueller matrices obtained
Existing techniques to do polarimetric endoscopic characterizations

## Slide 8
Summary
How to find Mueller matrix of a sample through an optical fiber ?
2) Polarimetric characteristics measurements of calibrated samples
  Polarimetric characteristics measurement of a waveplate
  Linear phase retardance measurement
  Linear diattenuation measurement
Polarimetric characteristics measurement of a linear retarder associated with
a linear diattenuator
4) Alternative technique to avoid fiber contribution
5) Conclusion

## Slide 9
Summary
How to find Mueller matrix of a sample through an optical fiber ?
2) Polarimetric characteristics measurements of calibrated samples
  Polarimetric characteristics measurement of a waveplate
  Linear phase retardance measurement
  Linear diattenuation measurement
Polarimetric characteristics measurement of a linear retarder associated with
a linear diattenuator
4) Alternative technique to avoid fiber contribution
5) Conclusion

## Slide 10
How to find Mueller matrix of a sample through an optical fiber ? - Fiber Mueller matrix
3
2
4
1 : Injection lens
2 : Collimation lens
3 : Single mode fiber
4 : Detection with photodiode and data processing
PSG
CW laser diode @638nm
1
Measured matrix of
a single mode fiber
Lu & Chipman
decomposition [6]
PSA
Depolarization
Diattenuation
1 | 0 | 0.001 | 0.004
0.007 | -0.712 | 0.696 | -0.093
0.003 | -0.289 | -0.173 | 0.925
-0.003 | 0.632 | 0.691 | 0.337
0,04%
0,67%
Retardance
Total : 2,46 rads
Linear : 1,23 rads
[6] : “Interpretation of Mueller matrices based on polar decomposition”. S.Y. Lu and R. A. Chipman, JOSA A, Vol. 13, Issue 5, pp. 1106-1113 (1996)

## Slide 11
How to find Mueller matrix of a sample through an optical fiber ? - Mathematical explanation
x
y
Fast
Slow
x
y
Fast
Slow
θ
Waveplate with δ retardance
Oriented waveplate with δ retardance

## Slide 12
How to find Mueller matrix of a sample through an optical fiber ? - Mathematical explanation
Input side
Output side
« Endoscopic » optical fiber
Fast
Fast
Slow
Slow

## Slide 13
1 : Polarization insensitive beamsplitter cube
2 : Injection lens
3 : Single mode fiber
4 : Collimation lens
5 : Switchable mirror
6 : Sample
7 : Mirror
8 : Detection with photodiode and data processing
How to find Mueller matrix of a sample through an optical fiber ? - Experimental setup
How to deduce the polarimetric response of sample through fiber ?
Two measurements
Fiber
Fiber + sample
1
2
3
4
7
8
6
CW laser diode @638nm
PSG
PSA
5

## Slide 14
How to find Mueller matrix of a sample through an optical fiber ? - Mathematical explanation
Mirror
Switchable mirror
Sample
Beam exiting
the fiber
Towards fiber
and analysis
Mirror
Switchable mirror
Sample
Beam exiting
the fiber
Towards fiber
and analysis
Switchable mirror ON
Switchable mirror OFF
θ1
δ
Forward
Backward
Forward
Backward

## Slide 15
Experimental validation
with different types of samples :
Linear retarders (fixed of variable)
Diattenuators
Association of components
How to find Mueller matrix of a sample through an optical fiber ?
1
2
3
4
7
8
1 : Polarization insensitive beamsplitter cube
2 : Injection lens
3 : Single mode fiber
4 : Collimation lens
5 : Switchable mirror
6 : Sample
7 : Mirror
8 : Detection with photodiode and data processing
6
CW laser diode @638nm
PSG
PSA
5

## Slide 16
Summary
How to find Mueller matrix of a sample through an optical fiber ?
2) Polarimetric characteristics measurements of calibrated samples
  Polarimetric characteristics measurement of a waveplate
  Linear phase retardance measurement
  Linear diattenuation measurement
Polarimetric characteristics measurement of a linear retarder associated with
a linear diattenuator
4) Alternative technique to avoid fiber contribution
5) Conclusion

## Slide 17
Measurements of calibrated samples : Polarimetric characteristics of a waveplate
λ/8 @633nm waveplate
Mirror
Switchable mirror
Beam exiting the fiber
Towards fiber and analysis
x
y
z
rotation
λ/8 @633nm waveplate
y
z
x
λ/8 retardance @638nm :
  Single pass : 44,64°
  Double pass : 89,29°

## Slide 18
Measurements of calibrated samples : Polarimetric characteristics of a waveplate
@638nm
Chart: Orientation angle of waveplate [degrees]
| 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
Measured orientation | 0 | 10.5 | 21 | 31 | 42 | 50.5 | 61 | 71 | 81 | 91
 | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
 | 91.0431 | 89.6107 | 90.2066 | 90.0862 | 89.5992 | 89.158 | 90.2467 | 88.9002 | 88.2986 | 90.6134
 | 89.2955 | 89.2955 | 89.2955 | 89.2955 | 89.2955 | 89.2955 | 89.2955 | 89.2955 | 89.2955 | 89.2955

## Slide 19
Summary
How to find Mueller matrix of a sample through an optical fiber ?
2) Polarimetric characteristics measurements of calibrated samples
  Polarimetric characteristics measurement of a waveplate
  Linear phase retardance measurement
  Linear diattenuation measurement
Polarimetric characteristics measurement of a linear retarder associated with
a linear diattenuator
4) Alternative technique to avoid fiber contribution
5) Conclusion

## Slide 20
Chart: Babinet-Soleil compensator retardance [degrees]
| 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180
Measured | 3.0382 | 11.5796 | 19.4904 | 28.3185 | 38.4076 | 47.121 | 58.9873 | 68.2739 | 78.1911 | 87.5924 | 97.7389 | 108.6306 | 118.8344 | 129.8408 | 137.6943 | 147.4968 | 158.1019 | 166.414 | 177.3057
Expected | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180
Measurements of calibrated samples : Linear phase retardance of a Babinet-Soleil compensator
Babinet-Soleil Compensator :
tunable linear retarder
Mirror
Switchable mirror
Beam exiting the fiber
Towards fiber and analysis

## Slide 21
Summary
How to find Mueller matrix of a sample through an optical fiber ?
2) Polarimetric characteristics measurements of calibrated samples
  Polarimetric characteristics measurement of a waveplate
  Linear phase retardance measurement
  Linear diattenuation measurement
Polarimetric characteristics measurement of a linear retarder associated with
a linear diattenuator
4) Alternative technique to avoid fiber contribution
5) Conclusion

## Slide 22
Measurements of calibrated samples : Linear diattenuation measurement
Chart: Angle of incidence α on glass plate [degrees]
| 0 | 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | 45 | 50 | 55 | 60 | 65 | 70 | 75 | 80 | 85
Measured | 0.008 | 0.009 | 0.005 | 0.019 | 0.03 | 0.05 | 0.07 | 0.1 | 0.138 | 0.187 | 0.234 | 0.297 | 0.38 | 0.45 | 0.57 | 0.66 | 0.76 | 0.88
Tilted glass plate with α angle :
tunable linear diattenuator
Mirror
Switchable mirror
Beam exiting the fiber
Towards fiber and analysis
α

## Slide 23
Summary
How to find Mueller matrix of a sample through an optical fiber ?
2) Polarimetric characteristics measurements of calibrated samples
  Polarimetric characteristics measurement of a waveplate
  Linear phase retardance measurement
  Linear diattenuation measurement
Polarimetric characteristics measurement of a linear retarder associated with
a linear diattenuator
4) Alternative technique to avoid fiber contribution
5) Conclusion

## Slide 24
Experimental validation : Association of components
Chart: Babinet-Soleil compensator retardance [degrees]
| 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180
Expected linear retardance | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180
Measured linear retardance | 3.2102 | 8.4841 | 22.9873 | 35.4268 | 33.4204 | 41.9045 | 54.6879 | 66.9554 | 73.949 | 88.8535 | 93.2675 | 103.8726 | 114.1338 | 123.2484 | 134.1401 | 150.7643 | 153.0573 | 160.5096 | 173.6943
Expected linear diattenuation | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18 | 0.18
Measured linear diattenuation | 0.182 | 0.181 | 0.184 | 0.182 | 0.183 | 0.185 | 0.183 | 0.179 | 0.182 | 0.184 | 0.182 | 0.185 | 0.185 | 0.176 | 0.183 | 0.179 | 0.18 | 0.186 | 0.18
Babinet-Soleil Compensator :
tunable linear retarder : FAST AXIS SET AT 0°
Mirror
Switchable mirror
Beam exiting the fiber
Towards fiber and analysis
Tilted glass plate with α angle :
Fixed linear diattenuator (≈17%)
α

## Slide 25
Chart: Babinet-Soleil compensator retardance [degrees]
| 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180
Measured linear retardance | 3.7261 | 6.1911 | 16.2803 | 26.4841 | 37.6624 | 45 | 56.1783 | 64.4331 | 74.8662 | 85.6433 | 97.5096 | 106.5669 | 115.5669 | 124.2803 | 137.2357 | 143.828 | 153.9745 | 166.0701 | 173.121
Expected linear retardance | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180
Measured linear diattenuation | 0.34 | 0.354 | 0.339 | 0.339 | 0.33 | 0.317 | 0.312 | 0.293 | 0.2789 | 0.259 | 0.233 | 0.215 | 0.19 | 0.17 | 0.134 | 0.12 | 0.092 | 0.052 | 0.031
Measured circular diattenuation | 0 | 0.024 | 0.056 | 0.083 | 0.116 | 0.131 | 0.166 | 0.183 | 0.209 | 0.237 | 0.254 | 0.278 | 0.296 | 0.302 | 0.324 | 0.346 | 0.351 | 0.356 | 0.36
Babinet-Soleil Compensator :
tunable linear retarder : FAST AXIS SET AT 45°
Mirror
Switchable mirror
Beam exiting the fiber
Towards fiber and analysis
Tilted glass plate with α angle :
Fixed linear diattenuator (≈35%)
α
Experimental validation : Association of components

## Slide 26
Summary
How to find Mueller matrix of a sample through an optical fiber ?
2) Polarimetric characteristics measurements of calibrated samples
  Polarimetric characteristics measurement of a waveplate
  Linear phase retardance measurement
  Linear diattenuation measurement
Polarimetric characteristics measurement of a linear retarder associated with
a linear diattenuator
4) Alternative technique to avoid fiber contribution
5) Conclusion

## Slide 27
Alternative solution to the switchable mirror ?
1
2
3
4
6
7
5
CW laser diode @660nm
PSG
PSA
CW laser diode @638nm
7
638nm : characterization of fiber
660nm : characterization of fiber + sample
638nm
660nm
8
9
8
6
Challenge : deduce the linear retardance of fiber @660nm from 638nm
1 : Polarization insensitive beamsplitter cube
2 : Injection lens
3 : Single mode fiber
4 : Collimation lens
5 : Sample
6 : Mirror
7 : Detection with photodiode and data processing
8 : Dichroic mirror (45°)
9 : Dichroic mirror (straight)
638nm
660nm

## Slide 28
Conclusion
Conclusion
Perspectives
Capability of our method to overcome the fiber contribution
Several polarimetric characteristics of samples are accessible :
Rotation and retardance induced by linear retarders
Linear diattenuation
Circular diattenuation
Depolarization measurement of biological samples
Implementation of the chromatic method

## Slide 29
We are grateful to the french ANR for its financial support
to this work, through the IMULE project
Thanks to the workshop organizers &
thank you for your attention